TRANSPORT VALIDATION PROTOCOL

Medical Device Quality Management System
Document ID: PROTOCOL-TRA-XXXX | Rev: 1.0

| Document Type: | Validation Protocol | Effective Date: | [DD/MM/YYYY] |
| --- | --- | --- | --- |
| Document ID: | PROTOCOL-TRA-XXXX | Review Date: | [DD/MM/YYYY] |

1. Overview

This protocol describes the approach for validating the distribution and transportation process of medical devices to demonstrate its capability to consistently protect product profile, sterile packaging integrity, and functionality against environmental and physical stresses during transit. The validation study establishes the methodology, critical transport parameters, sampling requirements, acceptance criteria, and evaluation methods necessary to assess the effectiveness and consistency of the shipping configuration. Successful execution of this protocol will provide documented evidence to support routine implementation of the validated transportation pathways.

2. Objective

The objective of this validation protocol is to establish documented evidence that the shipping configurations and transit routes, when executed under defined operating conditions, consistently meet the predetermined acceptance criteria. The study is intended to verify the effectiveness, repeatability, and reproducibility of the protective packaging system through the evaluation of real-world or simulated shipping hazards, environmental monitoring results, and physical package integrity. The validation also confirms that the transport configuration remains capable of maintaining device safety and efficacy during routine global distribution operations.

3. Scope

This protocol applies to the transport validation of finished medical devices manufactured by the company. The validation includes the assessment of temperature/humidity profiles during transit, physical resistance to shock, vibration, and compression, and evaluation of the post-transport package integrity and device functionality against established acceptance criteria. The study covers all designated shipping configurations, worst-case distribution routes (including domestic and international transit lines), and third-party logistics (3PL) handling.

4. Responsibility

The roles of each department in planning, executing, reviewing, and approving the validation study are defined below:

| Department | Responsibility |
| --- | --- |
| Quality Assurance (QA) | Prepare, review, and approve the validation protocol and report. Ensure that the transport validation is executed as per the approved protocol, review the collected data, investigate deviations, and approve the final validation outcome. |
| Regulatory Affairs (RA) | Ensure compliance with regional and international distribution regulations (e.g., ISO 13485, FDA 21 CFR 820). Review and verify that the transport validation data supports product licensing and market placement requirements. |
| Production / Logistics | Ensure the medical devices and shipping configurations scheduled for validation are prepared and made available according to the protocol. Execute the physical shipping and handling steps under routine distribution operations. |
| Engineering / Packaging | Ensure that the required configurations (cartons, pallets, shippers) are available and that monitoring equipment (calibrated data loggers, shock indicators) are functional and installed properly inside test shippers. |

5. Frequency

The transport validation shall be conducted as a 3-run study under routine distribution conditions to verify consistency. Revalidation shall be performed whenever significant changes are made to the distribution process, including changes to the secondary/tertiary packaging materials, shipping box dimensions, choice of logistics service provider (courier), target transit routes with harsher environmental profiles, or when post-market surveillance data indicates an increased trend of transit-related product damage. Revalidation shall also be performed following significant facility or procedural modifications affecting the validated product configurations.

6. Critical Process Parameters (CPPs) & Critical Quality Attributes (CQAs)

The conditions maintained during transport are considered Critical Process Parameters (CPPs) because they directly influence the preservation of sterile barrier systems, device physical structural integrity, and overall device performance. During validation, parameters such as shock, vibration, temperature limits, and humidity variations shall be monitored and justified by evaluating the package's ability to consistently achieve the required physical and microbiological acceptance criteria under routine operating conditions.

Validation of these parameters is performed to demonstrate that exposure to environmental extremes lower than or higher than the approved product tolerances may compromise sterile integrity, fracture sensitive device components, or degrade materials, resulting in a failure to provide adequate patient safety and clinical efficacy.

7. Procedure

7.1. Pre-validation Activities

Verify the calibration status of all environmental data loggers, shock trackers, and physical inspection equipment.

Document the baseline physical and aesthetic integrity of the product's primary (sterile pouch/blister), secondary (shelf box), and tertiary (shipping shipper) packaging.

Map and document the selected test distribution routes, including planned transfer hubs and total estimated transit duration.

7.2. Execution of Transport Validation

Assemble the sample medical devices inside their approved commercial packaging configuration.

Position calibrated data loggers inside the worst-case locations within the shipping carton (e.g., geometric center, top corner near outer walls) as specified in the placement plan.

Ship the test configurations via designated courier services through the predefined routes, accounting for seasonal temperature fluctuations (summer/winter worst-case challenges).

7.3. Environmental and Transit Monitoring

Upon arrival at the designated destination or back at the receiving facility, download and extract data from all integrated temperature and shock loggers.

Continuously log any physical incidents, handling delays, or severe weather conditions encountered during the shipping execution.

7.4. Evaluation of Results

Conduct visual inspections of all tertiary and secondary cartons for signs of crushing, punctures, or water damage.

Perform sterile barrier integrity testing (e.g., dye penetration test per ASTM F1929 or bubble emission per ASTM F2096) on the primary packaging to verify seal security.

Perform functional performance testing on the enclosed medical devices to confirm they operate within specification limits post-transit.

8. Acceptance Criteria

Packaging Integrity: No visible structural breaches, major collapses, or punctures on the outer shipping cartons that affect the internal contents.

Sterile Barrier Security: Primary device packaging must show 100% integrity with no channel leaks, delamination, or tears when tested via dye penetration or bubble emission methods.

Environmental Thresholds: Internal shipper temperature and humidity levels must remain within the product’s specified storage limits (15°C to 30°C unless otherwise specified) throughout the entire transit duration.

Device Functionality: The medical devices must pass all post-transport physical, mechanical, and electrical functional verification checklists with zero failures.

9. Conclusion

(To be completed after protocol execution) The transport validation protocol has been executed in accordance with the specified guidelines. The gathered data has been thoroughly analyzed against the predetermined acceptance criteria to conclude whether the transport process is validated or requires adjustments.

10. Reference

ISO 13485: Medical devices — Quality management systems — Requirements for regulatory purposes.

ASTM D4169: Standard Practice for Performance Testing of Shipping Containers and Systems.

ISTA 2A/3A: International Safe Transit Association Pack Testing Standards.

FDA 21 CFR Part 820.150: Distribution requirements for medical devices.

11. Annexures (Data Recording Tables)

11.1. Annexure I: Transport Validation Data Recording Sheet

| Run No. | Consignment / AWB No. | Dispatch Date & Time | Arrival Date & Time | Origin Location | Destination Location | Total Transit Duration (Hrs) | Done By (Sign/Date) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 01 |  |  |  |  |  |  |  |
| 02 |  |  |  |  |  |  |  |
| 03 |  |  |  |  |  |  |  |

Comments / Observations during shipment loading:
___________________________________________________________________________________________________
___________________________________________________________________________________________________

11.2. Annexure II: Transport Environmental Parameter Recording Sheet

| Run No. | Logger ID / Serial No. | Set Temp. Range (°C) | Max Temp. Recorded (°C) | Min Temp. Recorded (°C) | Max Humidity Recorded (%) | MKT (°C)* | Spike/Excursion Details (If Any) | Verified By (Sign/Date) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 01 |  |  |  |  |  |  |  |  |
| 02 |  |  |  |  |  |  |  |  |
| 03 |  |  |  |  |  |  |  |  |

*MKT = Mean Kinetic Temperature
Data File Download Confirmation Name & Location: ______________________________________________

11.3. Annexure III: Packaging Integrity Data Sheet (Pre- and Post-Transport)

| Run No. | Sample ID / Box No. | Outer Shipper Integrity (Pass/Fail) | Inner Shelf Box Condition (Complies/No) | Primary Sterile Barrier Test (Dye/Bubble) | Seal Integrity Result (Pass/Fail) | Device Functionality Test (Pass/Fail) | Evaluated By (Sign/Date) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 01 | Box 1 |  |  |  |  |  |  |
| 01 | Box 2 |  |  |  |  |  |  |
| 02 | Box 1 |  |  |  |  |  |  |
| 02 | Box 2 |  |  |  |  |  |  |
| 03 | Box 1 |  |  |  |  |  |  |
| 03 | Box 2 |  |  |  |  |  |  |

11.4. Annexure IV: Transit Route Mapping / Data Logger Placement Plan

[Diagrammatic Representation / Sketch Layout Space for Logger Placement Within Carton]
+--------------------------------------------+
|  [Logger Pos 1: Top Corner Wall]           |
|                                            |
|                 [Logger Pos 2: Center]     |
|                                            |
|  [Logger Pos 3: Bottom Base]               |
+--------------------------------------------+

| Shipper Config ID | Total Shipped Qty | No. of Internal Loggers Placed | Specific Logger Orientations / Fixation Method | Prepared By (Sign/Date) |
| --- | --- | --- | --- | --- |

11.5. Annexure V: Observation and Deviation Record

| Deviation No. | Identified During (Run # / Stage) | Description of Deviation / Environmental Excursion | Initial Impact Analysis | Corrective Action Drafted (CAPA Ref #) | QA Approval (Sign/Date) |
| --- | --- | --- | --- | --- | --- |

11.6. Annexure VI: Validation Summary and Conclusion Sheet

| Assessment Parameter | Summary of Findings Across All 3 Runs | Compliance Status (Yes / No) |
| --- | --- | --- |
| Physical Outer Protection |  |  |
| Sterile Barrier Preservation |  |  |
| Thermal / Environmental Control |  |  |
| Device Performance Verification |  |  |

Final Conclusion:
[  ] The Transport Validation meets all criteria defined in this protocol. The transport configuration is considered validated.
[  ] The Transport Validation did not meet acceptance criteria. Action required.

Reviewed By (QA): ________________________   Date: ______________

11.7. Transport Validation Performance Record

| Activity | Name / Title | Department | Signature | Date |
| --- | --- | --- | --- | --- |
| Prepared By |  | Validation Engineer / Logistics |  |  |
| Reviewed By |  | Head of Production / Operations |  |  |
| Approved By |  | Head of Quality Assurance |  |  |