[MANUFACTURER NAME]

Ethylene Oxide (EO) Sterilization Validation Protocol

For [Product Family Name]

| Document Detail | Information |
| --- | --- |
| Document No. | [Document Number] |
| Revision No. | [Revision Number] |
| Effective Date | [DD-MM-YYYY] |
| Protocol Type | [Initial Validation / Revalidation] |
| Product Family | [Product Family Name] |
| Sterilization Method | Ethylene Oxide Gas |
| Prepared By | [Regulatory Affairs / Quality Assurance] |
| Reviewed By | [QA Head / Production Head / Microbiology] |
| Approved By | [Head QA / Authorized Signatory] |

# Content

Content	2

1. Protocol Approval Sheet	4

2. Revision History	4

3. Objective	4

4. Scope	5

5. Responsibility	5

6. Standards for Regulatory Requirement	6

7. List of Documents	6

8. Pre-Requisite	8

8.1 Equipment Qualification	8

8.2 Calibration	8

8.3 Product Qualification	8

8.4 Product and Packaging Material Evaluation	9

8.5 Product Grouping	9

8.6 Selection of Family Representative	10

9. Process and Validation Approach	11

9.1 Process Validation Procedure	11

9.2 Bioburden Assessment	11

9.3 Selection of Process Challenge Device (PCD)	12

9.4 BI or PCD Placement in the Product Load	12

9.5 Temperature and Humidity Sensor in Product Load	13

9.6 Load Temperature	13

9.7 Identification of the Worst-Case Location or Cold Spot	13

9.8 Process Qualification Runs	14

a) Half Cycle Method	14

b) Bioburden / BI Method	16

9.9 Maintenance of Validation	17

9.10 Requalification	18

10. Sampling Method and Test Method	20

11. List of Tests to be Executed	20

12. Process Parameter Check Steps	22

12.1 Environmental Preconditioning	22

12.2 Initial Evacuation	22

12.3 Humidification	22

12.4 Gas Injection and Gas Dwell	22

12.5 Post-Exposure Gas Purge and Air In-Bleed	23

12.6 Heated Aeration	23

Appendix 1. Simulation of Anticipated Process Conditions	24

Appendix 2. Determination of Bioburden	24

Appendix 3. Glossary	25

Appendix 4. Reference Standards and Guidance Documents	25

Final Validation Conclusion Template	27

# 1. Protocol Approval Sheet

This Ethylene Oxide (EO/ETO) Sterilization Validation Protocol is prepared to define the validation approach, responsibilities, acceptance criteria, test methods, documentation requirements, and final reporting requirements for EO sterilization of [Product Family Name] manufactured by [Manufacturer Name]. The protocol shall be reviewed and approved before execution. Any change in the approved protocol shall be controlled through the applicable document control and change control procedure.

| Prepared By |  | Reviewed By | Approved By |
| --- | --- | --- | --- |
| Protocol Information | Details |  |  |
| Protocol Title | Ethylene Oxide Sterilization Validation Protocol |  |  |
| Company | [Manufacturer Name] |  |  |
| Document No. | [Document Number] |  |  |
| Revision | [Revision Number] |  |  |
| Effective Date | [DD-MM-YYYY] |  |  |
| Validation Type | [Initial Validation / Revalidation] |  |  |
| Product Family | [Product Family Name] |  |  |
| Sterilization Site | [In-house / Approved Contract Sterilizer (External)] |  |  |
| Sterility Assurance Level | SAL 10⁻⁶ |  |  |
| Applicable Cycle | [To be filled from approved sterilizer cycle] |  |  |

# 2. Revision History

| Revision No. | Effective Date | Description of Change | Prepared By | Approved By |
| --- | --- | --- | --- | --- |
| 00 |  | New protocol prepared for EO sterilization validation of [Product Family Name]. |  |  |
| 01 |  | To be updated if there is change in product, packaging, load pattern, cycle parameter, sterilization site, or regulatory requirement. |  |  |

# 3. Objective

The objective of this protocol is to establish and document the validation requirements for the Ethylene Oxide sterilization process used for [Manufacturer Name] [Product Family Name]. The validation shall demonstrate that the selected EO sterilization cycle is capable of consistently achieving the required Sterility Assurance Level (SAL) of 10⁻⁶ while ensuring that the product, packaging, material properties, product performance, and EO residual levels remain within acceptable limits.

| Sr. No. | Objective |
| --- | --- |
| 1 | To validate the EO sterilization process for selected [Product Family Name] products. |
| 2 | To demonstrate that the process consistently achieves SAL 10⁻⁶. |
| 3 | To identify the worst-case product, product family representative, and worst-case load location. |
| 4 | To confirm EO gas penetration into the most difficult-to-sterilize location. |
| 5 | To establish BI / PCD placement and acceptance criteria. |
| 6 | To verify product bioburden before sterilization and ensure the validated cycle remains suitable. |
| 7 | To confirm product and packaging compatibility with EO sterilization. |
| 8 | To verify EO, ECH, and EG residuals, as applicable, are within acceptable limits. |
| 9 | To define routine process monitoring, product release, and requalification requirements. |

# 4. Scope

This protocol is applicable to EO sterilization validation of sterile [Product Family Name] manufactured by [Manufacturer Name]. The protocol applies to finished products packed in their final sterile barrier packaging and intended to be supplied sterile.

[If sterilization is performed by an external contract sterilizer: state that the manufacturer remains the legal manufacturer and is responsible for qualification and approval of the contract sterilizer, review and acceptance of all validation data and cycle records, final product release decision, maintenance of validation records in the Device Master File, and full regulatory compliance; and that the contract sterilizer shall execute the EO cycle strictly as per the approved protocol and associated quality agreement. If sterilization is performed in-house, state the in-house facility and equipment covered instead.]

| Product Category | Example Products Covered / To be Added |
| --- | --- |
| [Product category 1] | [List representative product/variant names covered by this protocol, where EO sterilized.] |
| [Product category 2] | [List representative product/variant names covered by this protocol, where EO sterilized.] |

# 5. Responsibility

| Department / Person | Responsibility |
| --- | --- |
| Regulatory Affairs | Ensure the protocol is aligned with applicable regulatory and standard requirements. Maintain the protocol and final report in the DMF / technical documentation and support regulatory submissions. |
| Quality Assurance | Review and approve protocol, monitor validation execution, review sterilization records and test reports, evaluate deviations, approve final validation report, and authorize release criteria. |
| Production | Provide representative products, batch manufacturing records, cleaning records, product identification, and load quantity details. |
| Packaging Department | Ensure final packaging is performed using approved packaging material and validated sealing parameters. Provide packaging records and inspection status. |
| Quality Control / Microbiology | Coordinate bioburden testing, BI/PCD incubation, sterility testing, and review microbiological data. |
| [Sterilization Facility — in-house or contract sterilizer] | Execute the EO sterilization cycle strictly as per the approved protocol [and quality agreement, if external]. Provide sterilization cycle records, chamber printouts, equipment IQ/OQ/PQ qualification records, calibration certificates, BI/PCD incubation reports, deviation notifications, and any other data requested by QA. Notify QA immediately of any cycle deviation, equipment failure, or non-conformance. Maintain product under segregated, controlled conditions before and after sterilization. |
| Warehouse / Stores | Ensure segregation, storage, handling, and identification of pre-sterile and sterile products. |
| Authorized Signatory | Provide final approval of the validation protocol and final validation report. |

# 6. Standards for Regulatory Requirement

The following standards, guidance documents, and regulatory requirements shall be considered for EO sterilization validation. The latest applicable revision shall be verified before final approval of this protocol.

| Sr. No. | Standard / Regulation | Application in Protocol |
| --- | --- | --- |
| 1 | ISO 10012 | Control and calibration of measuring equipment used during EO validation. |
| 2 | ISO 10993-7 | EO and ECH residual limits, measurement, and product release after EO sterilization. |
| 3 | ISO 11135 | Main standard for EO sterilization development, validation, routine control, and revalidation. |
| 4 | ISO 11138-1 | General requirements for biological indicators used in sterilization validation. |
| 5 | ISO 11138-2 | Biological indicators specifically for EO sterilization processes. |
| 6 | ISO 11140-1 | General requirements for chemical indicators used to confirm EO exposure. |
| 7 | ISO 11737-1 | Determination and characterization of product bioburden before sterilization. |
| 8 | ISO 11737-2 | Sterility testing during validation and maintenance of sterilization process. |
| 9 | ISO 13485 | QMS requirements for validation, records, traceability, supplier control, and document control. |
| 10 | ISO 11607-1 | Requirements for sterile barrier materials, sterile barrier systems, and packaging systems. |
| 11 | ISO 11607-2 | Validation of forming, sealing, and assembly processes for sterile packaging. |
| 12 | ISO 14971 | Risk management for worst-case selection, EO residual risk, change impact, and revalidation need. |
| 13 | AAMI TIR28 | Product adoption and process equivalence for EO sterilization. |
| 14 | [Applicable national medical device regulation] | Regulatory requirements for sterile medical devices and implant documentation. |
| 15 | [Applicable regional regulation, e.g. EU MDR 2017/745, if applicable] | Technical documentation and validation evidence for sterile devices intended for that market. |

# 7. List of Documents

The following documents shall be available before, during, or after execution of the EO sterilization validation protocol. Document numbers and revision status shall be verified during protocol execution.

| Sr. No. | Document Name | Document No. / Revision / Attachment |
| --- | --- | --- |
| 1 | Approved EO sterilization validation protocol |  |
| 2 | Product master list and product codes |  |
| 3 | Approved product drawings and specifications |  |
| 4 | Bill of material and material certificates |  |
| 5 | Manufacturing batch records |  |
| 6 | Cleaning records |  |
| 7 | Packaging specification and sterile barrier material certificate |  |
| 8 | Packaging process validation report |  |
| 9 | Label artwork and label specification |  |
| 10 | Sterilization facility approval / supplier qualification records and, if external, signed quality agreement defining responsibilities, cycle parameters, record submission, deviation notification, and regulatory compliance obligations |  |
| 11 | Sterilizer IQ/OQ/PQ or equipment qualification records |  |
| 12 | Calibration certificates for critical instruments |  |
| 13 | EO sterilization cycle record / chamber printout |  |
| 14 | Load configuration diagram |  |
| 15 | BI/PCD placement diagram |  |
| 16 | Bioburden test report |  |
| 17 | BI incubation report |  |
| 18 | Sterility test report |  |
| 19 | EO residual test report |  |
| 20 | Packaging integrity test report |  |
| 21 | Post-sterilization product inspection report |  |
| 22 | Deviation / non-conformance report, if any |  |
| 23 | Final EO sterilization validation report |  |

# 8. Pre-Requisite

The following pre-requisites shall be completed and verified before execution of EO sterilization validation. Validation shall not be initiated until all critical pre-requisites are available or justified by QA.

## 8.1 Equipment Qualification

[If sterilization is performed externally: describe obtaining and reviewing the contract sterilizer's current IQ/OQ/PQ records for the sterilizer chamber, preconditioning room, aeration chamber, vacuum system, humidification system, EO gas injection system, air/nitrogen purge system, and monitoring/recording systems before initiating validation, and confirming acceptability as part of supplier qualification. If in-house: describe the equivalent in-house equipment qualification requirements.]

| Qualification Activity | Purpose | Required Evidence |
| --- | --- | --- |
| Installation Qualification (IQ) | To verify that equipment and utilities are installed as per approved specification. | Approved IQ report, facility/utility verification, equipment identification, installation checklist. |
| Operational Qualification (OQ) | To verify that the sterilizer operates within defined operating limits. | Approved OQ report, alarm checks, control functions, empty chamber runs, parameter verification. |
| Performance Qualification (PQ) | To verify consistent performance during defined product load conditions. | Approved PQ report or current validation records, load mapping, and related evidence. |

## 8.2 Calibration

All instruments used for monitoring, measuring, or recording critical EO sterilization parameters shall have valid calibration status before validation execution. Calibration certificates for all critical instruments shall be obtained and verified before validation commencement and at each subsequent sterilization batch. Instruments found out of calibration shall not be used unless an impact assessment is completed and approved by QA.

| Instrument / Device | Calibration Requirement | Acceptance Criteria |
| --- | --- | --- |
| Temperature sensors / data loggers | Valid calibration certificate traceable to national/international standards. | Calibration valid on date of use. |
| Humidity sensors | Valid calibration certificate. | Calibration valid and within tolerance. |
| Pressure gauges / transducers | Valid calibration certificate. | Calibration valid and within tolerance. |
| EO gas monitoring / dosing system | Calibration or verification as per sterilizer procedure. | System accepted for use. |
| Timer / recorder | Verification or calibration record. | Accurate and valid. |
| Weighing balance | Valid calibration certificate. | Calibration valid and within tolerance. |

## 8.3 Product Qualification

The product selected for EO sterilization validation shall be representative of routine production. Products shall be manufactured, cleaned, inspected, and packed as per approved procedures. The validation load shall represent the final commercial configuration wherever possible.

| Requirement | Status / Details (Yes / No / NA) |
| --- | --- |
| Product manufactured as per approved process |  |
| Product passed final inspection |  |
| Cleaning and drying completed |  |
| Final packaging completed |  |
| Batch / lot number available |  |
| Label and product identification verified |  |
| Manufacturing batch record completed |  |
| Product selected as family representative |  |

## 8.4 Product and Packaging Material Evaluation

Product and packaging materials shall be evaluated for suitability with EO sterilization. The evaluation shall include EO penetration ability, material compatibility, EO residual retention, sterile barrier integrity, label legibility, and any effect on product performance or safety.

| Material / Component | EO Sterilization Consideration |
| --- | --- |
| [Material type 1, e.g. Titanium alloy] | [Compatibility consideration, e.g. verify surface finish, dimensions, identification marking, absence of discoloration or damage.] |
| [Material type 2, e.g. Stainless Steel] | [Compatibility consideration, e.g. verify absence of corrosion, discoloration, staining.] |
| [Material type 3, e.g. polymer component] | [Polymer materials may retain EO more than metals; consider as worst-case for residual testing.] |

## 8.5 Product Grouping

Products may be grouped into a product family when they have similar materials, packaging configuration, manufacturing process, sterilization method, and sterile claim. Grouping shall be scientifically justified and documented. Products having significantly different material, complex geometry, dense packaging, or higher residual retention shall be evaluated separately or selected as the worst-case family representative.

| Grouping Criteria | Details to be Evaluated |
| --- | --- |
| Material similarity | [List material types applicable to this product family.] |
| Packaging similarity | Same or equivalent sterile barrier system, pouch size, carton, and sealing configuration. |
| Manufacturing similarity | Same cleaning, drying, inspection, and packing process. |
| Product design similarity | Similar size, geometry, holes, threads, cavities, or surface features. |
| Sterilization similarity | Same EO cycle, load pattern, aeration condition, and release criteria. |
| Sterile claim similarity | Products supplied sterile and intended for single-use implantation or surgical use. |

## 8.6 Selection of Family Representative

The family representative shall be selected as the product or configuration that is most difficult to sterilize and/or most difficult to aerate. Polymer-containing products may be considered worst-case for EO residual testing, while products with complex geometry, cavities, threads, dense packaging, or maximum load density may be considered worst-case for EO penetration.

| Selection Criteria | Worst-Case Basis |
| --- | --- |
| Largest product size | Largest implant or maximum packed volume may create a dense load. |
| Complex geometry | Holes, threads, cavities, porous surfaces, assembled parts, and restricted gas pathways. |
| Material composition | [Materials that may retain EO more than others, e.g. polymer vs. metallic components.] |
| Packaging configuration | Multiple layers, large carton, dense shipper, or restricted gas flow. |
| Load density | Maximum number of products per carton or maximum load configuration. |
| Aeration requirement | Product expected to require the longest aeration time. |
| Historical or test data | Previous residual, BI, or temperature mapping data indicating difficult locations. |

# 9. Process and Validation Approach

EO sterilization validation shall be performed using a documented and controlled approach. The validation shall include physical qualification and microbiological qualification of the sterilization process. Critical parameters such as preconditioning, temperature, humidity, vacuum, pressure, EO concentration, gas dwell time, purge, air in-bleed, and aeration shall be monitored and recorded.

| Process Stage | Purpose |
| --- | --- |
| Environmental preconditioning | To condition product load to required temperature and humidity before EO exposure. |
| Initial evacuation | To remove air from chamber and support EO penetration. |
| Humidification | To provide adequate moisture for microbial inactivation. |
| EO gas injection | To introduce defined EO gas concentration into chamber. |
| Gas dwell / exposure | To allow EO contact with products, packaging, and microorganisms. |
| Post-exposure purge | To remove EO from chamber after exposure. |
| Aeration | To reduce EO residuals to acceptable levels. |
| Microbiological qualification | To demonstrate BI/PCD inactivation and sterilization effectiveness. |
| Residual testing | To confirm product safety after aeration. |
| Packaging and product checks | To confirm product and sterile barrier system remain acceptable. |

## 9.1 Process Validation Procedure

The validation shall be executed according to this approved protocol. The product load shall include the selected family representative and shall be packed in final commercial packaging. Biological indicators or PCDs shall be placed at identified worst-case locations. Temperature and humidity sensors shall be placed at representative locations in the product load and chamber. Each validation run shall be reviewed for compliance with defined cycle parameters and acceptance criteria.

1. Verify pre-requisites and approval status of this protocol.

2. Prepare and identify the validation load as per defined load configuration.

3. Place BI/PCD at the approved worst-case locations.

4. Place temperature and humidity sensors in the load as per mapping plan.

5. Execute the EO sterilization cycle as per approved cycle parameters.

6. Collect BI/PCD and samples after sterilization and send for testing.

7. Review cycle records, sensor data, BI results, bioburden data, residual results, packaging integrity, and product inspection.

8. Record deviations, if any, and complete investigation before final approval.

9. Prepare and approve the final EO sterilization validation report.

## 9.2 Bioburden Assessment

Bioburden assessment shall be performed before sterilization to determine the viable microbial population present on the finished product. Samples shall be taken from representative pre-sterilized products after final cleaning and packaging, unless the approved method specifies otherwise. The test shall be performed using a validated or suitable method in accordance with ISO 11737-1.

| Parameter | Requirement |
| --- | --- |
| Standard | ISO 11737-1 |
| Sample type | Finished product before sterilization |
| Sampling stage | After final cleaning and packaging, before EO sterilization |
| Method | Membrane filtration / rinse method / extraction method, as applicable |
| Result expression | CFU per device or CFU per sample |
| Acceptance criteria | Within established alert/action limits or justified for validation |
| Review requirement | Any abnormal or high bioburden result shall be investigated. |

## 9.3 Selection of Process Challenge Device (PCD)

A Process Challenge Device (PCD) shall be selected or designed to present a challenge equal to or greater than the most difficult-to-sterilize location in the product load. The PCD may be an internal PCD, external PCD, or product PCD containing a biological indicator. The organism typically used for EO sterilization biological indicators is Bacillus atrophaeus.

| Parameter | Details |
| --- | --- |
| PCD type | Internal PCD / external PCD / product PCD |
| BI organism | Bacillus atrophaeus spores |
| BI population | Typically 10⁶ spores as per supplier certificate |
| BI resistance | As per BI certificate of analysis |
| Placement | Worst-case EO penetration location |
| Acceptance criteria | No growth after incubation for exposed BI/PCD; growth in positive control; no growth in negative control |

## 9.4 BI or PCD Placement in the Product Load

BI or PCD placement shall be based on the load configuration, product geometry, package density, expected EO penetration difficulty, and temperature distribution. Placement locations shall be identified in a load diagram and attached to the protocol.

| Placement Location | Justification |
| --- | --- |
| Center of load | Usually represents lower gas penetration and higher load density. |
| Middle carton or shipper | Represents dense package configuration. |
| Inside product package | Confirms EO penetration through sterile barrier system. |
| Product cavity / thread / internal feature | Represents complex geometry and restricted gas access. |
| Top, bottom, front, rear, and corner locations | Confirms distribution across chamber and load. |
| Identified cold spot | Confirms sterilization effectiveness at lowest temperature area. |

## 9.5 Temperature and Humidity Sensor in Product Load

Temperature and humidity sensors shall be placed at representative and worst-case locations in the product load. Sensor placement shall support identification of cold spots, humidity distribution, and confirmation that load conditions remain within validated limits.

| Sensor Type | Recommended Placement | Acceptance Criteria |
| --- | --- | --- |
| Temperature sensor | Center of load, top, bottom, front, rear, corner, and suspected cold spot. | Within validated cycle range. |
| Humidity sensor | Product load center and representative humidification location. | Within validated cycle range. |
| Chamber sensor | Sterilizer chamber monitoring point. | Within validated cycle range. |
| Product/location sensor | Near or inside representative product pack, where feasible. | Within validated cycle range. |

## 9.6 Load Temperature

The product load shall reach the required minimum temperature before EO gas exposure. EO exposure shall not be initiated until the load temperature is within the specified range. Load temperature data shall be reviewed to verify that all monitored locations meet the acceptance criteria.

| Parameter | Acceptance Criteria / Requirement |
| --- | --- |
| Minimum load temperature | As per approved validated cycle. |
| Maximum load temperature | As per approved validated cycle. |
| Temperature variation | Within approved operating limits. |
| Cold spot identification | Location with lowest measured temperature shall be identified and evaluated. |
| Cycle start condition | EO exposure shall begin only after required conditioning is achieved. |

## 9.7 Identification of the Worst-Case Location or Cold Spot

The worst-case location or cold spot shall be identified based on temperature mapping, humidity distribution, load density, packaging configuration, product design, and EO gas penetration. BI/PCD placement shall include the identified worst-case location.

| Basis | Worst-Case Location Consideration |
| --- | --- |
| Lowest temperature | Location showing the lowest temperature during mapping. |
| Dense packing | Center of carton, shipper, or full load. |
| Restricted gas flow | Inside carton, inner package, or tightly packed product. |
| Difficult EO penetration | Product cavity, internal thread, assembled area, or porous feature. |
| Historical validation data | Previously identified difficult locations or deviations. |

## 9.8 Process Qualification Runs

Process qualification runs shall demonstrate that the EO sterilization cycle is effective, reproducible, and suitable for routine use. The exact number of runs shall be defined based on the selected validation approach, applicable standard requirements, and sterilizer procedure.

| Run Type | Purpose | Expected Output |
| --- | --- | --- |
| Half-cycle run | To demonstrate BI inactivation at reduced EO exposure time. | All exposed BIs show no growth; controls acceptable. |
| Full-cycle run | To confirm routine cycle performance under defined load conditions. | Cycle parameters within limits; product and packaging acceptable. |
| Residual evaluation run | To evaluate EO residuals after defined aeration. | EO/ECH/EG results within applicable limits. |
| Repeat / confirmation run | To demonstrate reproducibility, if required. | Consistent results across runs. |

### a) Half Cycle Method

The half-cycle method shall be performed by exposing the product load to EO sterilization conditions using half of the routine EO exposure time, while maintaining other critical parameters such as temperature, humidity, vacuum, pressure, and EO concentration within specified limits. Successful inactivation of BIs during the half cycle provides assurance that the full cycle is capable of achieving the required sterility assurance level.

| Parameter | Requirement |
| --- | --- |
| EO exposure time | Half of the proposed routine exposure time. |
| BI/PCD placement | Worst-case locations identified in the product load. |
| BI organism | Bacillus atrophaeus. |
| BI result | No growth after incubation for exposed BIs. |
| Positive control | Growth observed. |
| Negative control | No growth observed. |
| Cycle parameters | Within defined range or justified deviation. |

Flowchart:

Start

↓  Review bioburden assessment data

↓  If bioburden characterization is not performed or bioburden is more than 100 CFU/device, select the half-cycle method

↓  Identify worst-case product, worst-case load configuration, and cold spot location

↓  Place Biological Indicators containing 10⁶ spores of Bacillus Atrophaeus in Process Challenge Devices and product samples

↓  Place PCDs, product samples, temperature sensors, and humidity sensors at the worst-case / cold spot locations in each pallet

↓  Load pallets into the EO sterilization chamber as per approved load pattern

↓  Run fractional cycle using reduced EO exposure time, for example ¼ or ⅙ of the full-cycle exposure time

↓  Aerate the product/load

↓  Remove product samples and PCDs from the load

↓  Perform sterility testing on product samples and Biological Indicators

↓  Perform bacteriostasis / fungistasis test on product samples

↓  Review fractional cycle results

↓  If product sterility samples show survival, repeat the fractional cycle with increased EO exposure time until no survival is observed in product samples and appropriate BI challenge is confirmed

↓  Run three consecutive half-cycle runs using ½ of the routine full-cycle EO exposure time

↓  Aerate the load after each half-cycle run

↓  Remove product samples and PCDs from each run

↓  Perform sterility testing on product samples and Biological Indicators

↓  Acceptance Criteria: Product samples and BIs shall show no growth after the half-cycle runs

↓  After successful completion of three half cycles, run one full cycle using routine EO exposure time

↓  Place additional product samples for required post-sterilization testing

↓  Aerate the load as per validated aeration condition

↓  Remove product samples and PCDs

↓  Perform the following tests on product samples:

- LAL test / bacterial endotoxin test, if applicable

- Product functionality test

- Packaging integrity / packaging functionality test

- Biocompatibility evaluation, if applicable

- Ethylene oxide residual testing

↓  Review all validation results

↓  Prepare EO sterilization validation report

↓  Approval by QA / RA / authorized person

↓  EO sterilization cycle considered validated for routine use

### b) Bioburden / BI Method

The Bioburden / BI method uses product bioburden data together with biological indicators of known population and resistance. Product bioburden shall be determined on representative samples before sterilization. The EO process shall be challenged using BI/PCDs at worst-case locations. The method supports demonstration that the sterilization process is adequate for the product bioburden and selected sterility assurance level.

| Parameter | Acceptance Criteria |
| --- | --- |
| Product bioburden | Within established or justified limits. |
| BI population and resistance | As per approved BI certificate. |
| BI result after EO exposure | No growth. |
| Product sterility test, if performed | No growth. |
| Method suitability | Pass / acceptable recovery and no inhibition. |
| Deviation | Investigated, justified, and approved before validation closure. |

Flowchart:

Start

↓  Perform bioburden assessment and characterization

↓  Confirm that bioburden level is consistent and less than or equal to 100 CFU/device

↓  Confirm that product bioburden is less resistant than or equal to the Biological Indicator challenge

↓  Select Bioburden / BI method

↓  Identify worst-case product, load configuration, and cold spot based on temperature and humidity distribution

↓  Place Biological Indicators with known spore population of Bacillus Atrophaeus in PCDs

↓  Place PCDs, temperature sensors, and humidity sensors at worst-case locations in the product load

↓  Perform five fractional cycles with graded EO exposure times, for example 0, 4, 8, 12, and 16 minutes

↓  Aerate the load after each fractional cycle

↓  Remove PCDs and Biological Indicators

↓  Perform sterility testing on Biological Indicators

↓  Calculate microbial inactivation rate / D-value using survival curve method or fraction-negative method

↓  Calculate required half-cycle and full-cycle EO exposure time

↓  Place product samples and PCDs in each pallet with temperature and humidity sensors

↓  Run three consecutive half-cycle runs using the calculated half-cycle exposure time

↓  Aerate the load after each half-cycle run

↓  Remove product samples and PCDs

↓  Perform sterility testing on product samples and Biological Indicators

↓  Perform bacteriostasis / fungistasis test on product samples

↓  Acceptance Criteria: Product samples and BIs shall show no growth after the half-cycle runs

↓  After successful completion of three half cycles, run one full cycle using routine EO exposure time

↓  Place additional product samples for required post-sterilization testing

↓  Aerate the load as per validated aeration condition

↓  Remove product samples and PCDs

↓  Perform the following tests on product samples:

- LAL test / bacterial endotoxin test, if applicable

- Product functionality test

- Packaging integrity / packaging functionality test

- Biocompatibility evaluation, if applicable

- Ethylene oxide residual testing

↓  Review all validation results

↓  Prepare EO sterilization validation report

↓  Approval by QA / RA / authorized person

↓  EO sterilization cycle considered validated for routine use

## 9.9 Maintenance of Validation

The validated EO sterilization process shall be maintained through routine process monitoring, periodic review, bioburden trend analysis, sterilization record review, and change control. Any adverse trend or deviation shall be evaluated to determine whether requalification or revalidation is required.

| Data to Review | Frequency / Trigger |
| --- | --- |
| Sterilization cycle records | Each batch. |
| BI / PCD results | Each validation run and routine batch, as applicable. |
| Chemical indicator results | Each batch, if used. |
| Bioburden trends | Periodic review as per SOP. |
| EO residual results | At validation and as per revalidation or change assessment. |
| Packaging integrity data | Validation, stability, and routine monitoring as applicable. |
| Deviations and CAPA | Each occurrence. |
| Change controls | Whenever change occurs. |
| Sterilization facility performance review (cycle record review, deviation history, calibration status, equipment qualification status, and quality agreement compliance) | Periodic review. |

## 9.10 Requalification

Requalification shall be performed periodically or whenever significant changes occur that may affect sterilization effectiveness, EO residuals, product safety, or packaging integrity. All changes shall be evaluated through change control and risk assessment before implementation.

| Trigger | Requalification Requirement |
| --- | --- |
| Change in product design | Impact assessment; requalification if sterilization difficulty or residual profile changes. |
| Change in material | Requalification required, especially for polymer or absorbent material changes. |
| Change in packaging material or size | Impact assessment; requalification if EO penetration or sterile barrier is affected. |
| Change in load configuration | Requalification required if maximum load, carton arrangement, or density changes. |
| Change in EO cycle parameters | Requalification required. |
| Change of sterilization facility, chamber, or site | Requalification required. |
| Increase in bioburden trend | Investigation and possible requalification. |
| BI failure or sterility failure | Immediate investigation and validation impact assessment. |
| EO residual failure | Investigation of aeration and possible cycle or aeration requalification. |
| Periodic review | As per company SOP or applicable standard requirement. |

# 10. Sampling Method and Test Method

Sampling shall be performed using representative products from the validation load. Samples shall be identified with product name, product code, batch number, sterilization batch number, sample quantity, and test requirement. Sample handling shall avoid mix-up or contamination.

| Test | Sample Type |  | Sample Stage | Method / Standard |
| --- | --- | --- | --- | --- |
| Bioburden | Finished product |  | Before EO sterilization | ISO 11737-1 / approved lab method |
| BI / PCD incubation | BI or PCD |  | After EO cycle | BI supplier IFU / approved microbiology method |
| Product sterility test | Sterilized product |  | Post-sterilization | ISO 11737-2 / approved lab method |
| EO residual | Worst-case product |  | After defined aeration | ISO 10993-7 / approved lab method |
| Packaging integrity | Sterilized packed product |  | Post-sterilization | ISO 11607 / approved internal method |
| Visual inspection | Sterilized product and package |  | Post-sterilization | Approved inspection procedure |
| Dimensional / functional check | Sterilized product |  | Post-sterilization | Approved product specification / drawing |
| Test |  | Suggested Sample Quantity / Selection |  |  |
| Bioburden |  | As per approved sampling plan and laboratory method. |  |  |
| Sterility |  | As per validated or approved sampling plan. |  |  |
| EO residual |  | Worst-case product and sample quantity as required by ISO 10993-7/laboratory. |  |  |
| Packaging integrity |  | Representative samples from sterilized load. |  |  |
| Product inspection |  | Representative samples from sterilized load or as per QA plan. |  |  |
| BI / PCD |  | Quantity and placement as per load size and worst-case location plan. |  |  |

# 11. List of Tests to be Executed

| Sr. No. | Test Name | Purpose | Acceptance Criteria |
| --- | --- | --- | --- |
| 1 | Equipment qualification review | To confirm sterilizer and associated equipment qualification. | IQ/OQ/PQ records available and approved. |
| 2 | Calibration verification | To confirm critical instruments are calibrated. | Valid calibration status on date of use. |
| 3 | Product inspection before sterilization | To confirm product suitability. | Product passes approved specification. |
| 4 | Packaging inspection before sterilization | To confirm correct sterile barrier packaging. | No damage, correct seal, correct label. |
| 5 | Bioburden test | To determine microbial load before sterilization. | Within established or justified limits. |
| 6 | BI / PCD test | To challenge EO sterilization cycle. | No growth in exposed BI/PCD; controls acceptable. |
| 7 | Temperature monitoring | To confirm cycle temperature and load temperature. | Within validated range. |
| 8 | Humidity monitoring | To confirm required humidity during conditioning/cycle. | Within validated range. |
| 9 | EO gas concentration / dosing | To confirm EO exposure. | Within validated range. |
| 10 | Pressure / vacuum monitoring | To confirm cycle vacuum and pressure profile. | Within validated range. |
| 11 | Product sterility test | To confirm no microbial growth, if performed. | No growth. |
| 12 | EO residual test | To confirm residual safety after aeration. | Within ISO 10993-7 limits. |
| 13 | Packaging integrity test | To confirm sterile barrier integrity. | Pass as per packaging method. |
| 14 | Visual inspection after sterilization | To detect product or packaging damage. | No unacceptable change. |
| 15 | Functional / dimensional check | To confirm product performance after EO. | Meets product specification. |
| 16 | Label verification | To confirm readability and traceability. | Label legible and correct. |

# 12. Process Parameter Check Steps

The EO sterilization cycle shall be monitored step by step. Actual values shall be recorded in the sterilizer cycle record or chamber printout and submitted to QA for review against approved acceptance criteria. The values below are placeholders and shall be replaced by approved cycle parameters confirmed with the sterilization facility before validation execution.

## 12.1 Environmental Preconditioning

The product load shall be placed in the environmental preconditioning area before sterilization. Preconditioning helps the product load reach the required temperature and humidity before EO exposure and supports effective microbial inactivation.

| Parameter | Target / Range | Actual | Acceptance |
| --- | --- | --- | --- |
| Temperature | ___ °C | ___ | Within range |
| Relative humidity | ___ % RH | ___ | Within range |
| Preconditioning time | ___ hours | ___ | Minimum time completed |
| Load condition | Final packed product | ___ | Correct load identified |

## 12.2 Initial Evacuation

After loading the sterilizer chamber, initial evacuation shall be performed to remove air from the chamber and product load. Proper air removal supports EO penetration into packaging and product surfaces.

| Parameter | Target / Range | Actual | Acceptance |
| --- | --- | --- | --- |
| Vacuum level | ___ mbar / kPa | ___ | Achieved |
| Number of vacuum pulses | ___ | ___ | As per cycle |
| Hold time | ___ minutes | ___ | As per cycle |
| Leak / pressure stability | As per procedure | ___ | Acceptable |

## 12.3 Humidification

Humidification shall be performed to achieve the required relative humidity inside the chamber and product load. Adequate humidity is an important factor in EO sterilization effectiveness.

| Parameter | Target / Range | Actual | Acceptance |
| --- | --- | --- | --- |
| Relative humidity | ___ % RH | ___ | Within range |
| Humidification time | ___ minutes | ___ | Completed |
| Steam / water vapor injection | As per cycle | ___ | Completed |
| Humidity distribution | As per sensor data | ___ | Acceptable |

## 12.4 Gas Injection and Gas Dwell

EO gas shall be injected into the sterilization chamber after achieving required vacuum, temperature, and humidity conditions. The load shall be exposed to EO gas for the defined dwell time. EO concentration, exposure time, temperature, humidity, and pressure shall remain within validated limits.

| Parameter | Target / Range | Actual | Acceptance |
| --- | --- | --- | --- |
| EO gas concentration | ___ mg/L | ___ | Within range |
| Gas injection time | ___ minutes | ___ | As per cycle |
| Exposure / dwell time | ___ hours | ___ | Completed |
| Chamber temperature | ___ °C | ___ | Within range |
| Chamber pressure | ___ mbar / kPa | ___ | Within range |

## 12.5 Post-Exposure Gas Purge and Air In-Bleed

After completion of EO exposure, the chamber shall be evacuated to remove EO gas. Repeated vacuum and air or nitrogen in-bleed cycles shall be performed to reduce EO concentration in the chamber and product load.

| Parameter | Target / Range | Actual | Acceptance |
| --- | --- | --- | --- |
| Number of purge cycles | ___ | ___ | As per cycle |
| Vacuum level | ___ | ___ | Achieved |
| Air / nitrogen in-bleed | As per cycle | ___ | Completed |
| Purge time | ___ minutes | ___ | Completed |
| Safety release condition | As per sterilizer procedure | ___ | Acceptable |

## 12.6 Heated Aeration

After EO exposure and gas purge, the product load shall be transferred to the aeration area or aeration chamber. Heated aeration shall be performed for the defined duration to reduce EO and related residuals to acceptable limits. EO residual testing shall be performed on the selected worst-case product after the defined aeration period.

| Parameter | Target / Range | Actual | Acceptance |
| --- | --- | --- | --- |
| Aeration temperature | ___ °C | ___ | Within range |
| Aeration time | ___ hours | ___ | Minimum time completed |
| Air exchange / ventilation | As per cycle | ___ | Acceptable |
| Product condition | Packed product | ___ | Correct identification |
| EO residual result | As per ISO 10993-7 | ___ | Pass |

# Appendix 1. Simulation of Anticipated Process Conditions

This appendix shall summarize the anticipated EO sterilization cycle profile agreed with the sterilization facility. Actual validated values shall be obtained from the approved sterilization cycle records and facility documentation, and incorporated into the final validation report. The information may be supported by a cycle graph showing temperature, humidity, pressure, EO exposure, purge, and aeration phases.

| Process Step | Expected Parameter / Placeholder |
| --- | --- |
| Preconditioning temperature | ___ °C |
| Preconditioning RH | ___ % RH |
| Preconditioning time | ___ hours |
| Initial vacuum | ___ mbar / kPa |
| Humidification RH | ___ % RH |
| EO gas concentration | ___ mg/L |
| Exposure time | ___ hours |
| Chamber temperature | ___ °C |
| Gas purge cycles | ___ |
| Aeration temperature | ___ °C |
| Aeration time | ___ hours |
| Load configuration | Attachment No. ___ |
| BI/PCD placement diagram | Attachment No. ___ |

# Appendix 2. Determination of Bioburden

Bioburden testing shall be performed on representative finished products before EO sterilization. Samples shall be collected from routine manufacturing batches after final cleaning and packaging, unless otherwise justified. Results shall be used to confirm the microbial load and support the sterilization validation approach.

| Parameter | Details |
| --- | --- |
| Product name | ___ |
| Product code | ___ |
| Batch number | ___ |
| Sample quantity | ___ |
| Test method | ___ |
| Extraction method | ___ |
| Culture media | ___ |
| Incubation condition | ___ |
| Result | ___ CFU/device |
| Alert limit | ___ |
| Action limit | ___ |
| Conclusion | Pass / Fail |

Conclusion statement: The bioburden result shall be reviewed against the established acceptance criteria. If the result is within the acceptable or justified limit, the product bioburden shall be considered suitable for EO sterilization validation. Any abnormal result shall be investigated before continuation or release.

# Appendix 3. Glossary

| Term | Meaning |
| --- | --- |
| EO / ETO | Ethylene Oxide. |
| SAL | Sterility Assurance Level. |
| SAL 10⁻⁶ | Probability of not more than one non-sterile item in one million sterilized items. |
| BI | Biological Indicator. |
| CI | Chemical Indicator. |
| PCD | Process Challenge Device. |
| RH | Relative Humidity. |
| CFU | Colony Forming Unit. |
| D-value | Time required under defined conditions to reduce microbial population by 90%. |
| Half Cycle | Cycle with reduced EO exposure time used to challenge sterilization process. |
| Full Cycle | Validated routine EO sterilization cycle. |
| Aeration | Process used to remove EO residues from product and packaging. |
| Bioburden | Number of viable microorganisms present on product before sterilization. |
| Cold Spot | Location with lowest temperature or most difficult sterilization condition. |
| EO Residual | EO remaining on or in product after sterilization. |
| ECH | Ethylene Chlorohydrin. |
| EG | Ethylene Glycol. |
| Sterile Barrier System | Packaging system that maintains sterility until point of use. |

# Appendix 4. Reference Standards and Guidance Documents

The protocol content is based on recognized sterilization, microbiological, packaging, quality management, and risk management standards. The latest applicable version of each standard shall be verified before final issue of the protocol.

| Sr. No. | Standard / Regulation | Application in Protocol |
| --- | --- | --- |
| 1 | ISO 10012 | Control and calibration of measuring equipment used during EO validation. |
| 2 | ISO 10993-7 | EO and ECH residual limits, measurement, and product release after EO sterilization. |
| 3 | ISO 11135 | Main standard for EO sterilization development, validation, routine control, and revalidation. |
| 4 | ISO 11138-1 | General requirements for biological indicators used in sterilization validation. |
| 5 | ISO 11138-2 | Biological indicators specifically for EO sterilization processes. |
| 6 | ISO 11140-1 | General requirements for chemical indicators used to confirm EO exposure. |
| 7 | ISO 11737-1 | Determination and characterization of product bioburden before sterilization. |
| 8 | ISO 11737-2 | Sterility testing during validation and maintenance of sterilization process. |
| 9 | ISO 13485 | QMS requirements for validation, records, traceability, supplier control, and document control. |
| 10 | ISO 11607-1 | Requirements for sterile barrier materials, sterile barrier systems, and packaging systems. |
| 11 | ISO 11607-2 | Validation of forming, sealing, and assembly processes for sterile packaging. |
| 12 | ISO 14971 | Risk management for worst-case selection, EO residual risk, change impact, and revalidation need. |
| 13 | AAMI TIR28 | Product adoption and process equivalence for EO sterilization. |
| 14 | [Applicable national medical device regulation] | Regulatory requirements for sterile medical devices and implant documentation. |
| 15 | [Applicable regional regulation, e.g. EU MDR 2017/745, if applicable] | Technical documentation and validation evidence for sterile devices intended for that market. |

# Final Validation Conclusion Template

Based on successful completion of equipment qualification review, calibration verification, product and packaging evaluation, bioburden assessment, BI/PCD challenge testing, process parameter monitoring, EO residual testing, packaging integrity testing, and final product inspection, the Ethylene Oxide sterilization process shall be considered validated for the selected [Manufacturer Name] [Product Family Name].

The validated EO sterilization cycle shall be used for routine sterilization only when all critical process parameters are maintained within approved limits and all acceptance criteria are met. [If sterilization is performed externally: state that the manufacturer shall review and approve all cycle records and test reports submitted by the sterilization facility before authorizing product release.] Any change in product design, material, packaging configuration, load pattern, sterilization cycle, sterilization facility, or regulatory requirement shall be evaluated through change control and may require requalification or revalidation.