[MANUFACTURER NAME]

Gamma Radiation Sterilization Validation Protocol

For [Product Family Name]

| Document Detail | Information |
| --- | --- |
| Document No. | [Document Number] |
| Revision No. | [Revision Number] |
| Effective Date | [DD-MM-YYYY] |
| Protocol Type | [Initial Validation / Revalidation / Dose Substantiation] |
| Product Family | [Product Family Name] |
| Sterilization Method | Gamma Irradiation (Cobalt-60) |
| Dose Setting Method | [VDmax25 / Method 1 / Method 2, as applicable] |
| Prepared By | [Regulatory Affairs / Quality Assurance] |
| Reviewed By | [QA Head / Production Head / Microbiology] |
| Approved By | [Head QA / Authorized Signatory] |

# Content

Content	1

1. Protocol Approval Sheet	1

2. Revision History	1

3. Objective	1

4. Scope	1

5. Responsibility	1

6. Standards for Regulatory Requirement	1

7. List of Documents	1

8. Pre-Requisite	1

8.1 Irradiator / Facility Qualification	1

8.2 Dosimetry System	1

8.3 Product Qualification	1

8.4 Product and Packaging Material Evaluation (Radiation Compatibility)	1

8.5 Product Grouping	1

8.6 Selection of Family Representative	1

9. Process and Validation Approach	1

9.1 Process Validation Procedure	1

9.2 Bioburden Assessment	1

9.3 Dose-Setting Methods (ISO 11137-2)	1

a) VDmax25 Method	1

b) Method 1 — Bioburden-Based Dose Setting	1

c) Method 2 — Dose Substantiation	1

9.4 Dose Mapping Study	1

9.5 Maximum Acceptable Dose and Material Compatibility	1

9.6 Sterility Testing	1

9.7 Maintenance of Validation and Routine Dosimetric Release	1

9.8 Requalification / Dose Audit	1

10. Sampling Method and Test Method	1

11. List of Tests to be Executed	1

12. Process Parameter Check Steps	1

12.1 Pre-Irradiation Checks	1

12.2 Irradiation Process Parameters	1

12.3 Post-Irradiation Checks	1

Appendix 1. Dose Mapping Study Summary	1

Appendix 2. Determination of Bioburden	1

Appendix 3. Glossary	1

Appendix 4. Reference Standards and Guidance Documents	1

Final Validation Conclusion Template	1

# 1. Protocol Approval Sheet

This Gamma Radiation Sterilization Validation Protocol is prepared to define the validation approach, responsibilities, acceptance criteria, dose-setting method, test methods, documentation requirements, and final reporting requirements for gamma sterilization of [Product Family Name] manufactured by [Manufacturer Name]. The protocol shall be reviewed and approved before execution. Any change in the approved protocol shall be controlled through the applicable document control and change control procedure.

| Prepared By |  | Reviewed By | Approved By |
| --- | --- | --- | --- |
| Protocol Information | Details |  |  |
| Protocol Title | Gamma Radiation Sterilization Validation Protocol |  |  |
| Company | [Manufacturer Name] |  |  |
| Document No. | [Document Number] |  |  |
| Revision | [Revision Number] |  |  |
| Effective Date | [DD-MM-YYYY] |  |  |
| Validation Type | [Initial Validation / Revalidation / Dose Substantiation] |  |  |
| Product Family | [Product Family Name] |  |  |
| Irradiation Facility | [In-house / Approved Contract Irradiator (External)] |  |  |
| Radiation Source | Cobalt-60 (Gamma) |  |  |
| Sterility Assurance Level | SAL 10⁻⁶ |  |  |
| Dose Setting Method | [VDmax25 / ISO 11137-2 Method 1 / Method 2] |  |  |
| Minimum Sterilization Dose | [To be established / __ kGy] |  |  |
| Maximum Acceptable Dose | [To be established / __ kGy] |  |  |

# 2. Revision History

| Revision No. | Effective Date | Description of Change | Prepared By | Approved By |
| --- | --- | --- | --- | --- |
| 00 |  | New protocol prepared for gamma sterilization validation of [Product Family Name]. |  |  |
| 01 |  | To be updated if there is change in product, packaging, load pattern, dose, irradiation site, or regulatory requirement. |  |  |

# 3. Objective

The objective of this protocol is to establish and document the validation requirements for the gamma radiation sterilization process used for [Manufacturer Name] [Product Family Name]. The validation shall demonstrate that the selected sterilization dose is capable of consistently achieving the required Sterility Assurance Level (SAL) of 10⁻⁶, that the minimum dose is sufficient to achieve sterility, and that the maximum dose does not adversely affect product safety, material properties, packaging integrity, or performance.

| Sr. No. | Objective |
| --- | --- |
| 1 | To establish the sterilization dose for [Product Family Name] in accordance with ISO 11137-2. |
| 2 | To demonstrate that the process consistently achieves SAL 10⁻⁶ at the established minimum dose. |
| 3 | To identify the worst-case product / product family representative for dose setting. |
| 4 | To perform dose mapping and identify minimum and maximum dose zones within the irradiation container. |
| 5 | To verify product bioburden before irradiation and confirm suitability of the selected dose-setting method. |
| 6 | To confirm product, material, and packaging compatibility with the maximum acceptable dose. |
| 7 | To verify sterility of irradiated product samples. |
| 8 | To define routine dosimetric release, process monitoring, and dose auditing (quarterly dose audit) requirements. |
| 9 | To establish requalification criteria for the validated dose. |

# 4. Scope

This protocol is applicable to gamma radiation sterilization validation of sterile [Product Family Name] manufactured by [Manufacturer Name]. The protocol applies to finished products packed in their final sterile barrier packaging and intended to be supplied sterile.

[If sterilization is performed by an external contract irradiator: state that the manufacturer remains the legal manufacturer and is responsible for qualification and approval of the contract irradiator, review and acceptance of all dosimetry data and irradiation records, final product release decision, maintenance of validation records in the Device Master File, and full regulatory compliance; and that the contract irradiator shall execute the irradiation cycle strictly as per the approved protocol and associated quality/technical agreement. If sterilization is performed in-house, state the in-house facility and irradiator covered instead.]

| Product Category | Example Products Covered / To be Added |
| --- | --- |
| [Product category 1] | [List representative product/variant names covered by this protocol, where gamma sterilized.] |
| [Product category 2] | [List representative product/variant names covered by this protocol, where gamma sterilized.] |

# 5. Responsibility

| Department / Person | Responsibility |
| --- | --- |
| Regulatory Affairs | Ensure the protocol is aligned with applicable regulatory and standard requirements. Maintain the protocol and final report in the DMF / technical documentation and support regulatory submissions. |
| Quality Assurance | Review and approve protocol, monitor validation execution, review dosimetry and irradiation records, evaluate deviations, approve final validation report, and authorize release criteria. |
| Production | Provide representative products, batch manufacturing records, cleaning records, product identification, and load quantity details. |
| Packaging Department | Ensure final packaging is performed using approved packaging material and validated sealing parameters. Provide packaging records and inspection status. |
| Quality Control / Microbiology | Coordinate bioburden testing, dosimetry analysis, sterility testing, and review microbiological data. |
| [Irradiation Facility — in-house or contract irradiator] | Execute the gamma irradiation cycle strictly as per the approved protocol [and quality/technical agreement, if external]. Provide dosimetry mapping data, irradiation process records, dosimeter calibration certificates, source strength/decay records, deviation notifications, and any other data requested by QA. Notify QA immediately of any dose deviation, equipment failure, or non-conformance. Maintain product under segregated, controlled conditions before and after irradiation. |
| Warehouse / Stores | Ensure segregation, storage, handling, and identification of pre-sterile and sterile products. |
| Authorized Signatory | Provide final approval of the validation protocol and final validation report. |

# 6. Standards for Regulatory Requirement

The following standards, guidance documents, and regulatory requirements shall be considered for gamma sterilization validation. The latest applicable revision shall be verified before final approval of this protocol.

| Sr. No. | Standard / Regulation | Application in Protocol |
| --- | --- | --- |
| 1 | ISO 11137-1 | Requirements for development, validation, and routine control of a radiation sterilization process for medical devices. |
| 2 | ISO 11137-2 | Establishing the sterilization dose (VDmax25, Method 1, Method 2). |
| 3 | ISO 11137-3 | Guidance on dosimetric aspects of development, validation, and routine control. |
| 4 | ISO/ASTM 51261 | Guide for selection and calibration of dosimetry systems for radiation processing. |
| 5 | ISO/ASTM 51649 / 51702 / 51204 (as applicable) | Practice for dosimetry in gamma irradiation facilities. |
| 6 | ISO 11737-1 | Determination and characterization of product bioburden before sterilization. |
| 7 | ISO 11737-2 | Sterility testing during validation and maintenance of sterilization process. |
| 8 | ISO 13485 | QMS requirements for validation, records, traceability, supplier control, and document control. |
| 9 | ISO 11607-1 | Requirements for sterile barrier materials, sterile barrier systems, and packaging systems. |
| 10 | ISO 11607-2 | Validation of forming, sealing, and assembly processes for sterile packaging. |
| 11 | ISO 10993-1 | Biological evaluation, including assessment of any radiation-induced material changes. |
| 12 | ISO 14971 | Risk management for worst-case selection, dose-related risk, change impact, and revalidation need. |
| 13 | AAMI TIR33 / TIR15 (as applicable) | Guidance on sterilization dose substantiation and audit methods. |
| 14 | [Applicable national medical device regulation] | Regulatory requirements for sterile medical devices and implant documentation. |
| 15 | [Applicable regional regulation, e.g. EU MDR 2017/745, if applicable] | Technical documentation and validation evidence for sterile devices intended for that market. |

# 7. List of Documents

The following documents shall be available before, during, or after execution of the gamma sterilization validation protocol. Document numbers and revision status shall be verified during protocol execution.

| Sr. No. | Document Name | Document No. / Revision / Attachment |
| --- | --- | --- |
| 1 | Approved gamma sterilization validation protocol |  |
| 2 | Product master list and product codes |  |
| 3 | Approved product drawings and specifications |  |
| 4 | Bill of material and material certificates |  |
| 5 | Manufacturing batch records |  |
| 6 | Cleaning records |  |
| 7 | Packaging specification and sterile barrier material certificate |  |
| 8 | Packaging process validation report |  |
| 9 | Label artwork and label specification |  |
| 10 | Irradiation facility approval / supplier qualification records and, if external, signed quality/technical agreement defining responsibilities, dose limits, record submission, deviation notification, and regulatory compliance obligations |  |
| 11 | Irradiator qualification records (source characterization, dose mapping capability) |  |
| 12 | Dosimeter calibration certificates |  |
| 13 | Dose mapping study report / product loading configuration |  |
| 14 | Bioburden test report |  |
| 15 | Verification dose irradiation and sterility test report (VDmax / incremental dose experiments, as applicable) |  |
| 16 | Routine dosimetry records for validation irradiation |  |
| 17 | Sterility test report |  |
| 18 | Packaging integrity test report |  |
| 19 | Material compatibility / functional test report |  |
| 20 | Post-irradiation product inspection report |  |
| 21 | Deviation / non-conformance report, if any |  |
| 22 | Final gamma sterilization validation report |  |

# 8. Pre-Requisite

The following pre-requisites shall be completed and verified before execution of gamma sterilization validation. Validation shall not be initiated until all critical pre-requisites are available or justified by QA.

## 8.1 Irradiator / Facility Qualification

[If sterilization is performed externally: describe obtaining and reviewing the contract irradiator's current facility qualification records, including source characterization, dose mapping capability of the irradiation container/tote, conveyor system qualification, and dosimetry system calibration, before initiating validation, and confirming acceptability as part of supplier qualification. If in-house: describe the equivalent in-house irradiator qualification requirements.]

| Qualification Activity | Purpose | Required Evidence |
| --- | --- | --- |
| Installation Qualification (IQ) | To verify that the irradiator and associated systems are installed as per specification. | Approved IQ report, source loading records, facility layout verification. |
| Operational Qualification (OQ) | To verify that the irradiator operates within defined operating limits (conveyor speed, source strength, timer). | Approved OQ report, control system checks, empty/loaded container trial runs. |
| Performance Qualification (PQ) / Dose Mapping | To verify dose distribution within the product/irradiation container under defined loading conditions. | Dose mapping study report identifying minimum and maximum dose zones (Dmin/Dmax) and dose uniformity ratio (DUR). |

## 8.2 Dosimetry System

All dosimeters used for measuring absorbed dose shall be calibrated and traceable to a national/international standard, and shall be used within their validated dose range and read using a calibrated instrument. Dosimeter batches shall have a valid calibration curve established against the reading instrument before use.

| Item | Requirement | Acceptance Criteria |
| --- | --- | --- |
| Routine dosimeters | Calibrated dosimeter batch with valid calibration curve. | Calibration valid on date of use; within validated dose range. |
| Reference/alanine dosimeters | Used for dose mapping and calibration verification, as applicable. | Traceable calibration; valid uncertainty budget. |
| Dosimeter reading instrument (e.g. spectrophotometer) | Valid calibration/verification. | Calibration valid on date of use. |
| Dose uniformity ratio (DUR) | Established from dose mapping study. | Documented Dmax/Dmin ratio for the product/container. |

## 8.3 Product Qualification

The product selected for gamma sterilization validation shall be representative of routine production. Products shall be manufactured, cleaned, inspected, and packed as per approved procedures. The validation load shall represent the final commercial configuration and maximum/minimum density loading pattern wherever possible.

| Requirement | Status / Details (Yes / No / NA) |
| --- | --- |
| Product manufactured as per approved process |  |
| Product passed final inspection |  |
| Cleaning and drying completed |  |
| Final packaging completed |  |
| Batch / lot number available |  |
| Label and product identification verified |  |
| Manufacturing batch record completed |  |
| Product selected as family representative |  |

## 8.4 Product and Packaging Material Evaluation (Radiation Compatibility)

Product and packaging materials shall be evaluated for suitability with gamma irradiation at the anticipated maximum dose. The evaluation shall include assessment of discoloration, embrittlement, oxidative degradation, mechanical property change, polymer chain scission/crosslinking, sterile barrier integrity, label legibility, and any effect on product performance or safety.

| Material / Component | Gamma Irradiation Consideration |
| --- | --- |
| [Material type 1, e.g. Titanium alloy / Stainless Steel] | [Metallic materials are generally radiation-stable; verify absence of discoloration, corrosion, or dimensional change.] |
| [Material type 2, e.g. UHMWPE or other polymer component] | [Polymers may be susceptible to chain scission or crosslinking at high doses; evaluate mechanical/functional properties post-irradiation, and consider antioxidant stabilization where applicable.] |
| [Material type 3, e.g. packaging film/pouch] | [Evaluate seal strength, packaging integrity, and any embrittlement or discoloration after irradiation.] |

## 8.5 Product Grouping

Products may be grouped into a product family when they have similar materials, packaging configuration, manufacturing process, sterilization method, and sterile claim. Grouping shall be scientifically justified and documented. Products having significantly different material composition (especially radiation-sensitive polymers), density, or bioburden profile shall be evaluated separately or selected as the worst-case family representative.

| Grouping Criteria | Details to be Evaluated |
| --- | --- |
| Material similarity | [List material types applicable to this product family, noting any radiation-sensitive polymers.] |
| Packaging similarity | Same or equivalent sterile barrier system, pouch size, carton, and sealing configuration. |
| Manufacturing similarity | Same cleaning, drying, inspection, and packing process. |
| Product design similarity | Similar size, geometry, mass, and density. |
| Sterilization similarity | Same dose range, loading pattern, and release criteria. |
| Sterile claim similarity | Products supplied sterile and intended for single-use implantation or surgical use. |

## 8.6 Selection of Family Representative

The family representative shall be selected as the product or configuration considered worst-case for both minimum dose sterility assurance (highest/most resistant bioburden, most protected microbial location) and maximum dose material tolerance (most radiation-sensitive material, e.g. certain polymers).

| Selection Criteria | Worst-Case Basis |
| --- | --- |
| Highest or most resistant bioburden | Product with highest historical/typical bioburden count or most radiation-resistant organisms. |
| Most radiation-sensitive material | Products containing polymers or materials most susceptible to dose-related degradation. |
| Product mass / density | Denser products may require dose mapping consideration for uniform dose delivery. |
| Packaging configuration | Configuration most likely to create dose variation across the load. |
| Historical or test data | Previous dose mapping, bioburden, or material compatibility data indicating worst case. |

# 9. Process and Validation Approach

Gamma sterilization validation shall be performed using a documented and controlled approach in accordance with ISO 11137-1 and ISO 11137-2. The validation shall include selection of the sterilization dose (via VDmax25, Method 1, or Method 2), dose mapping to characterize dose distribution, and confirmation that the maximum dose does not adversely affect the product.

| Process Stage | Purpose |
| --- | --- |
| Bioburden determination | To determine the viable microbial population on the product before irradiation. |
| Dose-setting method selection | To select an appropriate method (VDmax25, Method 1, or Method 2) to establish the minimum sterilization dose. |
| Verification dose experiment | To confirm that the tentative/selected minimum dose achieves the required SAL. |
| Dose mapping | To characterize dose distribution and identify minimum (Dmin) and maximum (Dmax) dose locations within the product/irradiation container. |
| Irradiation at established dose range | To irradiate product between the established minimum sterilization dose and the maximum acceptable dose. |
| Sterility testing | To confirm sterility of irradiated product samples, as applicable to the chosen method. |
| Material/product compatibility testing | To confirm product, material, and packaging remain acceptable after exposure to the maximum dose. |
| Routine dosimetric release and dose audit | To confirm ongoing suitability of the established dose through routine monitoring and periodic (e.g. quarterly) dose audits. |

## 9.1 Process Validation Procedure

The validation shall be executed according to this approved protocol using the selected dose-setting method. The product load shall include the selected family representative and shall be packed in final commercial packaging. Dosimeters shall be placed at identified minimum and maximum dose locations based on dose mapping.

1. Verify pre-requisites and approval status of this protocol.

2. Perform bioburden determination on representative pre-sterilization product samples.

3. Select the dose-setting method (VDmax25 / Method 1 / Method 2) based on bioburden level and applicability.

4. Perform the verification dose experiment / incremental dosing experiment as required by the selected method.

5. Perform sterility testing on irradiated product samples as required by the method.

6. Establish the minimum sterilization dose based on verification results.

7. Perform dose mapping of the product/irradiation container to identify Dmin and Dmax locations and confirm dose uniformity ratio.

8. Irradiate the validation load within the established dose range and monitor with routine dosimeters at Dmin and Dmax locations.

9. Perform post-irradiation testing: sterility (if applicable), material compatibility, packaging integrity, and product inspection.

10. Review all dosimetry and test data against acceptance criteria; record and investigate deviations, if any.

11. Prepare and approve the final gamma sterilization validation report.

## 9.2 Bioburden Assessment

Bioburden assessment shall be performed before sterilization to determine the viable microbial population present on the finished product. Samples shall be taken from representative pre-sterilization products after final cleaning and packaging. The test shall be performed using a validated or suitable method in accordance with ISO 11737-1.

| Parameter | Requirement |
| --- | --- |
| Standard | ISO 11737-1 |
| Sample type | Finished product before sterilization |
| Sampling stage | After final cleaning and packaging, before irradiation |
| Sampling frequency | Per product family / per defined sampling plan (e.g. periodic bioburden determination) |
| Method | Membrane filtration / rinse method / extraction method, as applicable |
| Result expression | CFU per device or CFU per sample |
| Acceptance criteria | Within established alert/action limits used for the selected dose-setting method |
| Review requirement | Any abnormal or high bioburden result shall be investigated and dose-setting method suitability reassessed. |

## 9.3 Dose-Setting Methods (ISO 11137-2)

The sterilization dose shall be established using one of the following recognized methods, selected based on product bioburden level, product family characteristics, and applicability.

| Method | Applicability | Basis |
| --- | --- | --- |
| VDmax25 (Method for Low Bioburden) | Products with average bioburden ≤ 1000 CFU/device (or lower, per selected VDmax table) and no unusually resistant organisms. | Fixed substantiation dose of 25 kGy (or 15 kGy for VDmax15) verified using a small sample set; delivers SAL 10⁻⁶ at a validated minimum dose derived from standard tables. |
| Method 1 (Bioburden-Based) | Products where bioburden-based dose setting is preferred, using incremental dosing and sterility testing. | Minimum dose established from product bioburden data using the incremental dosing and fraction-negative or survival-curve approach per ISO 11137-2 Annex B/C. |
| Method 2 (Fractional / Substantiation) | Products where a specific dose (e.g. an already-used dose for a similar product) is to be substantiated, or verification of an existing dose is required. | Incremental dose experiments at fractional doses of the proposed sterilization dose, followed by sterility testing to substantiate SAL 10⁻⁶ at the selected dose. |

### a) VDmax25 Method

The VDmax25 method is applicable for products with low and consistent bioburden. A verification dose is derived from a standard table based on the average bioburden of ten sample items, and a small number of samples is irradiated at this dose and tested for sterility to confirm that a substantiation dose of 25 kGy is capable of achieving SAL 10⁻⁶.

| Parameter | Requirement |
| --- | --- |
| Applicable standard | ISO 11137-2, Method for VDmax25 (or VDmax15, as applicable) |
| Bioburden sampling | 10 product samples for average bioburden determination |
| Verification dose | Determined from ISO 11137-2 VDmax25 table based on average bioburden |
| Verification sample size | As specified in the applicable VDmax table (commonly 10–20 samples) |
| Sterility test | Performed on samples irradiated at the verification dose |
| Acceptance criteria | No more than the allowed number of positives per the applicable VDmax table (typically zero or as specified) |
| Established minimum sterilization dose | 25 kGy (or 15 kGy, per selected table), once verification is successful |

Flowchart:

Start

↓  Select 10 representative product samples from routine production

↓  Determine average bioburden per ISO 11737-1

↓  Determine verification dose from the applicable ISO 11137-2 VDmax table based on average bioburden

↓  Select verification sample set as per the applicable VDmax table

↓  Irradiate verification samples at the determined verification dose

↓  Perform sterility testing on irradiated verification samples per ISO 11737-2

↓  Review results against the acceptance criteria in the applicable VDmax table

↓  If acceptance criteria are met, the substantiation dose (25 kGy or 15 kGy) is confirmed as the minimum sterilization dose

↓  If acceptance criteria are not met, investigate bioburden and repeat verification, or select an alternate dose-setting method

↓  Perform dose mapping of the product/irradiation container

↓  Irradiate validation load within established dose range (minimum sterilization dose to maximum acceptable dose)

↓  Perform material compatibility, packaging integrity, and product inspection on maximum-dose samples

↓  Review all validation results

↓  Prepare gamma sterilization validation report

↓  Approval by QA / RA / authorized person

↓  Gamma sterilization dose considered validated for routine use

### b) Method 1 — Bioburden-Based Dose Setting

Method 1 establishes the sterilization dose based on the average bioburden of the product using an incremental dosing approach and sterility testing, per ISO 11137-2 Annex B or C. This method may be used when a dose lower than 25 kGy is desired, or when VDmax25 verification does not succeed.

| Parameter | Requirement |
| --- | --- |
| Applicable standard | ISO 11137-2, Method 1 (Annex B — Method using fraction positive; Annex C — Method using bioburden information) |
| Bioburden sampling | As per selected annex requirements (typically 10 samples per production period across multiple periods) |
| Incremental dose experiment | Irradiation of samples at a series of incremental doses |
| Sterility test | Performed on samples at each incremental dose |
| Dose calculation | Minimum sterilization dose derived using the fraction-negative or bioburden-based calculation method |
| Acceptance criteria | As per the calculation method and applicable tables in ISO 11137-2 |

Flowchart:

Start

↓  Determine product bioburden across multiple representative production batches/periods as required by the selected annex

↓  Select incremental dose levels for the dose-setting experiment

↓  Irradiate product samples at each incremental dose

↓  Perform sterility testing on samples at each incremental dose per ISO 11737-2

↓  Calculate the minimum sterilization dose using the applicable ISO 11137-2 Annex B/C calculation method

↓  Confirm the calculated minimum dose meets the target SAL 10⁻⁶

↓  Perform dose mapping of the product/irradiation container

↓  Irradiate validation load within the established dose range (minimum sterilization dose to maximum acceptable dose)

↓  Perform sterility testing, material compatibility, packaging integrity, and product inspection on validation load samples

↓  Review all validation results

↓  Prepare gamma sterilization validation report

↓  Approval by QA / RA / authorized person

↓  Gamma sterilization dose considered validated for routine use

### c) Method 2 — Dose Substantiation

Method 2 is used to substantiate a specific proposed sterilization dose (for example, a dose already validated for an equivalent product, or a dose selected based on prior experience) using fractional dose experiments and sterility testing, per ISO 11137-2.

| Parameter | Requirement |
| --- | --- |
| Applicable standard | ISO 11137-2, Method 2 |
| Proposed dose | [State proposed dose to be substantiated, e.g. __ kGy] |
| Fractional dose experiment | Irradiation of samples at defined fractions of the proposed dose (e.g. 0, ⅓, and full fraction as applicable) |
| Sterility test | Performed on samples at each fractional dose |
| Acceptance criteria | Number of positives at each fraction shall meet ISO 11137-2 Method 2 tables to substantiate the proposed dose |

Flowchart:

Start

↓  Determine product bioburden per ISO 11737-1

↓  Define the proposed sterilization dose to be substantiated

↓  Select fractional dose levels as per ISO 11137-2 Method 2 requirements

↓  Irradiate product samples at each fractional dose

↓  Perform sterility testing on samples at each fractional dose per ISO 11737-2

↓  Compare the number of positive results at each fraction against ISO 11137-2 Method 2 acceptance tables

↓  Confirm whether the proposed dose is substantiated as the minimum sterilization dose

↓  If not substantiated, reassess bioburden, select an alternate dose, or select an alternate dose-setting method

↓  Perform dose mapping of the product/irradiation container

↓  Irradiate validation load within the established dose range (minimum sterilization dose to maximum acceptable dose)

↓  Perform sterility testing, material compatibility, packaging integrity, and product inspection on validation load samples

↓  Review all validation results

↓  Prepare gamma sterilization validation report

↓  Approval by QA / RA / authorized person

↓  Gamma sterilization dose considered validated for routine use

## 9.4 Dose Mapping Study

Dose mapping shall be performed to characterize the dose distribution within the product and irradiation container/tote under the defined loading pattern. The study shall identify the minimum dose (Dmin) and maximum dose (Dmax) locations and establish the dose uniformity ratio (DUR = Dmax / Dmin).

| Parameter | Requirement |
| --- | --- |
| Dosimeter placement | Representative locations throughout the product load / irradiation container, including corners, center, and product-dense areas |
| Number of dosimeters | Sufficient to characterize dose distribution as per irradiator qualification / dose mapping procedure |
| Reference product density / loading pattern | As defined for routine production loading configuration |
| Dmin location | Location receiving the lowest absorbed dose |
| Dmax location | Location receiving the highest absorbed dose |
| Dose uniformity ratio (DUR) | Calculated as Dmax / Dmin; used to set process parameters (conveyor speed / exposure time / source strength) |
| Acceptance criteria | Dmin ≥ established minimum sterilization dose; Dmax ≤ established maximum acceptable dose |

## 9.5 Maximum Acceptable Dose and Material Compatibility

The maximum acceptable dose shall be established based on material compatibility, product performance, and packaging integrity data. Product samples irradiated at or above the maximum dose (Dmax) identified during dose mapping shall be evaluated to confirm no adverse effect on safety or performance.

| Test / Evaluation | Purpose | Acceptance Criteria |
| --- | --- | --- |
| Visual inspection | Detect discoloration, embrittlement, or visible degradation. | No unacceptable change compared to control/unirradiated sample. |
| Dimensional check | Confirm no dimensional change due to irradiation. | Within approved product specification. |
| Mechanical / functional testing | Confirm product performance is maintained after maximum dose exposure. | Meets approved product specification. |
| Packaging integrity test | Confirm sterile barrier system integrity after irradiation. | Pass as per approved packaging test method (e.g. ISO 11607). |
| Biocompatibility re-evaluation, if applicable | Confirm no adverse biological effect due to radiation-induced material change. | As per ISO 10993-1 risk-based evaluation. |

## 9.6 Sterility Testing

Sterility testing shall be performed as required by the selected dose-setting method and on the final validation load, in accordance with ISO 11737-2.

| Parameter | Requirement |
| --- | --- |
| Standard | ISO 11737-2 |
| Sample source | Product samples irradiated at the minimum sterilization dose (dose-setting method) and/or the validation load |
| Test method | Direct inoculation or membrane filtration, as validated |
| Incubation | As per validated method and applicable culture media/conditions |
| Acceptance criteria | No growth (or number of positives within method-specific acceptance criteria for dose-setting experiments) |

## 9.7 Maintenance of Validation and Routine Dosimetric Release

Once the sterilization dose is established and validated, routine production irradiation shall be monitored using dosimeters placed at or near the previously established minimum and maximum dose locations. Each production irradiation batch shall be released based on dosimetry confirming the delivered dose is within the established dose range.

| Data to Review | Frequency / Trigger |
| --- | --- |
| Routine dosimetry records | Each irradiation batch. |
| Delivered dose vs. established range | Each irradiation batch (dosimetric release). |
| Bioburden trends | Periodic review as per SOP / quarterly dose audit. |
| Quarterly dose audit (per ISO 11137-2) | Every 3 months (or as defined by the applicable standard/SOP), to confirm the established dose remains adequate. |
| Packaging integrity data | Validation, stability, and routine monitoring as applicable. |
| Deviations and CAPA | Each occurrence. |
| Change controls | Whenever change occurs. |
| Irradiation facility performance review (dosimetry record review, deviation history, calibration status, and quality agreement compliance) | Periodic review. |

## 9.8 Requalification / Dose Audit

Requalification, dose substantiation, or re-establishment of the sterilization dose shall be performed periodically (e.g. quarterly dose audit per ISO 11137-2) or whenever significant changes occur that may affect bioburden, material compatibility, or dose distribution. All changes shall be evaluated through change control and risk assessment before implementation.

| Trigger | Requalification / Audit Requirement |
| --- | --- |
| Routine quarterly dose audit | Periodic bioburden determination and dose audit per ISO 11137-2 to confirm the established dose remains valid. |
| Change in product design | Impact assessment; re-establishment of dose if bioburden or material profile changes. |
| Change in material | Requalification required, especially for radiation-sensitive polymers or additives. |
| Change in packaging material or size | Impact assessment; dose mapping repeat if dose distribution is affected. |
| Change in loading pattern / product density | Dose mapping repeat required. |
| Change of irradiation facility, source, or site | Dose mapping and facility qualification review required. |
| Adverse bioburden trend | Investigation and possible dose re-establishment. |
| Sterility test failure | Immediate investigation and validation impact assessment. |
| Material compatibility failure at maximum dose | Investigation of maximum acceptable dose and possible re-evaluation. |
| Periodic review | As per company SOP or applicable standard requirement. |

# 10. Sampling Method and Test Method

Sampling shall be performed using representative products from the validation load and dose-setting experiments. Samples shall be identified with product name, product code, batch number, irradiation batch number, sample quantity, and test requirement. Sample handling shall avoid mix-up or contamination.

| Test | Sample Type | Sample Stage | Method / Standard |
| --- | --- | --- | --- |
| Bioburden | Finished product | Before irradiation | ISO 11737-1 / approved lab method |
| Dosimetry (Dmin/Dmax mapping) | Dosimeters placed in load/container | During irradiation | ISO/ASTM 51261 / 51649 / 51702 (as applicable) |
| Sterility test (dose-setting) | Irradiated verification/incremental dose samples | Post-irradiation | ISO 11737-2 |
| Sterility test (validation load) | Irradiated validation load samples | Post-irradiation | ISO 11737-2 |
| Material compatibility | Maximum-dose product samples | Post-irradiation | Approved internal method / relevant material standard |
| Packaging integrity | Irradiated packed product | Post-irradiation | ISO 11607 / approved internal method |
| Visual inspection | Irradiated product and package | Post-irradiation | Approved inspection procedure |
| Dimensional / functional check | Irradiated product | Post-irradiation | Approved product specification / drawing |

# 11. List of Tests to be Executed

| Sr. No. | Test Name | Purpose | Acceptance Criteria |
| --- | --- | --- | --- |
| 1 | Facility / irradiator qualification review | To confirm irradiator and dosimetry system qualification. | Qualification records available and approved. |
| 2 | Dosimeter calibration verification | To confirm dosimetry system is calibrated. | Valid calibration status on date of use. |
| 3 | Product inspection before irradiation | To confirm product suitability. | Product passes approved specification. |
| 4 | Packaging inspection before irradiation | To confirm correct sterile barrier packaging. | No damage, correct seal, correct label. |
| 5 | Bioburden test | To determine microbial load before irradiation. | Within established or justified limits for the selected dose-setting method. |
| 6 | Verification / incremental / fractional dose sterility test | To establish or substantiate the minimum sterilization dose. | Meets applicable ISO 11137-2 acceptance table for the selected method. |
| 7 | Dose mapping | To characterize dose distribution and confirm Dmin/Dmax locations. | Dmin ≥ minimum sterilization dose; Dmax ≤ maximum acceptable dose. |
| 8 | Routine dosimetry (validation load) | To confirm delivered dose is within the established range. | Within established dose range. |
| 9 | Sterility test (validation load) | To confirm sterility of irradiated product. | No growth. |
| 10 | Material compatibility test | To confirm no adverse effect at maximum dose. | Meets product/material specification. |
| 11 | Packaging integrity test | To confirm sterile barrier integrity. | Pass as per packaging method. |
| 12 | Visual inspection after irradiation | To detect product or packaging damage/discoloration. | No unacceptable change. |
| 13 | Functional / dimensional check | To confirm product performance after irradiation. | Meets product specification. |
| 14 | Label verification | To confirm readability and traceability. | Label legible and correct. |

# 12. Process Parameter Check Steps

The gamma irradiation cycle shall be monitored using placed dosimeters and facility process records. Actual values shall be recorded and submitted to QA for review against approved acceptance criteria. The values below are placeholders and shall be replaced by approved parameters confirmed with the irradiation facility before validation execution.

## 12.1 Pre-Irradiation Checks

| Parameter | Target / Range | Actual | Acceptance |
| --- | --- | --- | --- |
| Product/load identification | Correct batch and product family | ___ | Verified |
| Packaging condition | No damage, correct seal | ___ | Acceptable |
| Loading pattern | As per approved dose mapping configuration | ___ | Matches qualified pattern |
| Dosimeter placement | As per dose mapping plan (Dmin/Dmax locations) | ___ | Verified |

## 12.2 Irradiation Process Parameters

| Parameter | Target / Range | Actual | Acceptance |
| --- | --- | --- | --- |
| Minimum sterilization dose | ___ kGy | ___ | Delivered dose ≥ minimum |
| Maximum acceptable dose | ___ kGy | ___ | Delivered dose ≤ maximum |
| Conveyor speed / exposure (dwell) time | ___ | ___ | As per qualified cycle |
| Source activity / strength | ___ (as recorded by facility) | ___ | Within facility operating range |
| Product density / loading configuration | As per approved dose mapping configuration | ___ | Matches qualified pattern |

## 12.3 Post-Irradiation Checks

| Parameter | Target / Range | Actual | Acceptance |
| --- | --- | --- | --- |
| Dosimeter reading — Dmin location | ≥ established minimum dose | ___ | Pass |
| Dosimeter reading — Dmax location | ≤ established maximum dose | ___ | Pass |
| Product visual condition | No discoloration/damage | ___ | Acceptable |
| Packaging condition | No damage, seal intact | ___ | Acceptable |
| Product segregation and labeling (irradiated status) | Correctly identified | ___ | Verified |

# Appendix 1. Dose Mapping Study Summary

This appendix shall summarize the dose mapping study performed with the irradiation facility for the defined product loading configuration. Actual validated values shall be obtained from the approved dose mapping report and incorporated into the final validation report.

| Parameter | Value / Placeholder |
| --- | --- |
| Irradiation container / tote type | ___ |
| Product loading configuration | ___ |
| Number of dosimeters used | ___ |
| Minimum dose (Dmin) location and value | ___ |
| Maximum dose (Dmax) location and value | ___ |
| Dose uniformity ratio (DUR = Dmax/Dmin) | ___ |
| Established minimum sterilization dose | ___ kGy |
| Established maximum acceptable dose | ___ kGy |
| Load configuration diagram | Attachment No. ___ |
| Dosimeter placement diagram | Attachment No. ___ |

# Appendix 2. Determination of Bioburden

Bioburden testing shall be performed on representative finished products before gamma sterilization. Samples shall be collected from routine manufacturing batches after final cleaning and packaging, unless otherwise justified. Results shall be used to support the selected dose-setting method.

| Parameter | Details |
| --- | --- |
| Product name | ___ |
| Product code | ___ |
| Batch number | ___ |
| Sample quantity | ___ |
| Test method | ___ |
| Extraction method | ___ |
| Culture media | ___ |
| Incubation condition | ___ |
| Result | ___ CFU/device |
| Alert limit | ___ |
| Action limit | ___ |
| Conclusion | Pass / Fail |

Conclusion statement: The bioburden result shall be reviewed against the established acceptance criteria for the selected dose-setting method. If the result is within the acceptable or justified limit, the product bioburden shall be considered suitable for gamma sterilization dose setting. Any abnormal result shall be investigated before continuation or release.

# Appendix 3. Glossary

| Term | Meaning |
| --- | --- |
| Gamma Irradiation | Sterilization method using gamma radiation, typically from a Cobalt-60 source. |
| SAL | Sterility Assurance Level. |
| SAL 10⁻⁶ | Probability of not more than one non-sterile item in one million sterilized items. |
| kGy | Kilogray, the unit of absorbed radiation dose. |
| Dmin | Minimum absorbed dose location within the product/irradiation container. |
| Dmax | Maximum absorbed dose location within the product/irradiation container. |
| DUR | Dose Uniformity Ratio (Dmax / Dmin). |
| VDmax25 | Method for establishing a substantiation dose of 25 kGy for low-bioburden products, per ISO 11137-2. |
| Method 1 | Bioburden-based dose-setting method per ISO 11137-2. |
| Method 2 | Dose substantiation method using fractional dosing per ISO 11137-2. |
| Dose Mapping | Study to characterize dose distribution within a product/irradiation container. |
| Dosimeter | Device used to measure absorbed radiation dose. |
| Bioburden | Number of viable microorganisms present on product before sterilization. |
| Dosimetric Release | Product release based on confirmation that delivered dose is within the established range, without requiring sterility testing of each batch. |
| Sterile Barrier System | Packaging system that maintains sterility until point of use. |

# Appendix 4. Reference Standards and Guidance Documents

The protocol content is based on recognized sterilization, dosimetry, microbiological, packaging, quality management, and risk management standards. The latest applicable version of each standard shall be verified before final issue of the protocol.

| Sr. No. | Standard / Regulation | Application in Protocol |
| --- | --- | --- |
| 1 | ISO 11137-1 | Requirements for development, validation, and routine control of a radiation sterilization process for medical devices. |
| 2 | ISO 11137-2 | Establishing the sterilization dose (VDmax25, Method 1, Method 2). |
| 3 | ISO 11137-3 | Guidance on dosimetric aspects of development, validation, and routine control. |
| 4 | ISO/ASTM 51261 | Guide for selection and calibration of dosimetry systems for radiation processing. |
| 5 | ISO/ASTM 51649 / 51702 / 51204 (as applicable) | Practice for dosimetry in gamma irradiation facilities. |
| 6 | ISO 11737-1 | Determination and characterization of product bioburden before sterilization. |
| 7 | ISO 11737-2 | Sterility testing during validation and maintenance of sterilization process. |
| 8 | ISO 13485 | QMS requirements for validation, records, traceability, supplier control, and document control. |
| 9 | ISO 11607-1 | Requirements for sterile barrier materials, sterile barrier systems, and packaging systems. |
| 10 | ISO 11607-2 | Validation of forming, sealing, and assembly processes for sterile packaging. |
| 11 | ISO 10993-1 | Biological evaluation, including assessment of any radiation-induced material changes. |
| 12 | ISO 14971 | Risk management for worst-case selection, dose-related risk, change impact, and revalidation need. |
| 13 | AAMI TIR33 / TIR15 (as applicable) | Guidance on sterilization dose substantiation and audit methods. |
| 14 | [Applicable national medical device regulation] | Regulatory requirements for sterile medical devices and implant documentation. |
| 15 | [Applicable regional regulation, e.g. EU MDR 2017/745, if applicable] | Technical documentation and validation evidence for sterile devices intended for that market. |

# Final Validation Conclusion Template

Based on successful completion of facility/dosimetry qualification review, bioburden assessment, the selected dose-setting method (VDmax25 / Method 1 / Method 2), dose mapping, sterility testing, material compatibility testing, packaging integrity testing, and final product inspection, the gamma radiation sterilization process shall be considered validated for the selected [Manufacturer Name] [Product Family Name] at an established minimum sterilization dose of [___ kGy] and a maximum acceptable dose of [___ kGy].

The validated dose range shall be used for routine sterilization only when routine dosimetry confirms the delivered dose is within the established range for every irradiation batch. [If sterilization is performed externally: state that the manufacturer shall review and approve all dosimetry records and test reports submitted by the irradiation facility before authorizing product release.] Any change in product design, material, packaging configuration, loading pattern, irradiation facility, or regulatory requirement shall be evaluated through change control and may require dose mapping repeat, requalification, or revalidation. The established dose shall be confirmed through periodic dose audits (e.g. quarterly) in accordance with ISO 11137-2.